An Internal Review of the Radiotherapy Service at

University Hospitals Bristol NHS Foundation Trust

Division of Specialised Services

Final Report

June 2008

Table of Contents

1.	Executive Summary	3

2.	Introduction	4

3.	Terms of Reference of the Review	5

4.	Review Reports	5

4.1.	Staff input into assessment of clinical processes and facilities	5

4.2.	Assessment of root cause analyses undertaken on reportable incidents and on action plans	6

4.3.	Analysis of recent radiotherapy incidents to identify systemic issues and assessment against	6

4.3.1.	Analysis of Clinical Incidents	6

4.3.2.	Towards Safer Radiotherapy	7

4.4.	Inspection of compliance with the Ionising Radiation (Medical Exposures) Regulations, 2000, (IR(ME)R) by the Healthcare Commission	8

4.5.	Benchmark staffing, activity, capacity, waiting times	9

4.6.	Activity and capacity plan for 2008/09	10

5.	Summary of Actions Taken During the Review	11

6.	Final Review Action Plan	12

7.	Conclusion	15

Appendix A:  Review – rationale, terms of reference and membership of the team	i

Appendix B:  “Think Tank” invitation	ii

Appendix C:  “Think Tank” questionnaire	iii

Appendix D:  Radiotherapy Review Feedack Report, April 16 2008	iv

Appendix E:  Summary of Reportable Incidents and Actions Taken	vi

Appendix F:  Summary of Physics Incidents	viii

Appendix G:  Summary of Radiographic Incidents	ix

Appendix H:  Summary of Audit on Compliance with Towards Safer Radiotherapy	xi

Appendix I:  Benchmarking data	xv

- Executive Summary

A review of the Radiotherapy Service at University Hospitals Bristol NHS Foundation Trust was undertaken in response to the four reportable incidents under the Ionising Radiation (Medical Exposures) Regulations, 2000, (IR(ME)R) regulations between June 2007 and March 2008, and concerns raised by staff over workload in the department.

The Review team undertook a detailed analysis of Clinical Governance and capacity planning arrangements, and used staff involvement and benchmarking to assess the radiotherapy service’s ability to deliver safe and effective services.  These work areas are summarised below:

- Assessment of the root cause analyses of the four reportable incidents showed them to be almost entirely separate with one overlapping feature for two of the incidents.  The agreed action plans from each reportable incident were reported to the Trust’s Clinical Risk Advisory Committee in April 2008. The actions will be completed according to plan.

- The radiotherapy service was also audited against the standards in the recent publication Towards Safer Radiotherapy.  The audit found that, as with many services across the country, the Bristol service is mostly compliant and has plans in place to reach full compliance.  An analysis of recent clinical incidents in the Radiotherapy Physics Unit shows that the profile of incidents is not exceptional over the last two years, and a similar analysis of radiographic incidents demonstrated a range of error types.  Where a theme was apparent, an action plan has been developed and is being implemented.  The review concluded that the profile of reported incidents was in line with range expected in a standard radiotherapy department.

- The Healthcare Commission carried out an inspection of the department’s compliance with the (IR(ME)R) regulations on 17th June 2008.  The Inspectors reported that the radiotherapy service is compliant with the regulations and signed off the two completed action plans arising from reportable incidents. The inspection team highlighted areas of good practice but require some minor adjustments, mostly to documentation.  An action plan has been developed.

- A demand and capacity review was carried out on the period up to Taunton opening their new Beacon Oncology Centre in May 2009.  An activity plan was developed to meet the demand and implementation is underway.

- Meetings with various staff groups and grades highlighted areas of concern within the radiotherapy service.  All radiotherapy staff and Heads of Department at BHOC were invited to the Radiotherapy “Think Tank” event and encouraged to complete a short questionnaire to inform the agenda.  The Radiotherapy “Think Tank” event engaged over 80 staff at an evening workshop and produced an array of ideas and solutions which were recorded and themed into 8 work packages. All radiotherapy staff have been encouraged to volunteer to be involved with implementing the solutions and, with their help, many improvements are already complete or underway.

- Finally, benchmarking with other radiotherapy centres has demonstrated that the practices and procedures employed in BHOC are in line with those of the other three centres.

A summary action plan has been developed to deal with the issues identified during the review.  Overall however, this review was able to demonstrate that the radiotherapy service at University Hospitals Bristol NHS Foundation Trust is safe and effective.

The review provided an opportunity to improve the service until the Taunton Centre opens in May 2009 but also identified the need to work more closely with corporate and commissioning leads to define and deliver a shared strategy for future service provision.

- Introduction

The Division of Specialised Services places paramount importance on the safety of our patients and the wellbeing of our staff.

The Trust is aware that the years preceding the opening of the new Beacon Oncology Centre in Taunton in May 2009 represent a challenging time for the radiotherapy service, as explained in the Medium Term Plan for Non-Surgical Oncology, 2007.  The long-awaited and welcome expansion of medical workforce in Bristol and Taunton has revealed unmet need, particularly in Somerset.  This increased demand, coupled with the imperative of achieving the 31 and 62 day cancer targets and the underlying growth in cancer incidence, has furnished us the ever growing challenge of fitting more activity into the same five linear accelerators.

The department has taken measures to offset this impact, through extending workings hours from 8am to 8pm, putting a request to purchase and operate a sixth linear accelerator through the Trust capital programme, and appointing staff to fixed-term posts to deal with the bulge in demand pre-Taunton.

In spite of this, at the start of 2008 concerns were raised that four reportable incidents had taken place within six months and the staff were finding it difficult to manage the impact of the growing demand.  Even though this number of reportable incidents was not unduly high in comparison with other radiotherapy centres, it was higher than normal for Bristol

On consultation with Trust Executive Directors, the Division decided to conduct a formal internal review of the radiotherapy service.  A multi-disciplinary Review Team was formed from within the Division, which set out to investigate whether the service was safe and effective, and what could be done to make improvements for the safety and benefit of staff and patients, both up to the Beacon’s opening date and beyond.

The following report is the result of four months’ close scrutiny, hard work and dedication on the part of the Review Team and the staff of the radiotherapy service as a whole.  Clearly we found things that we could improve on, and many changes have already taken place or are underway within the department.  On the whole though, what we found was a group of staff who carry out their serious and difficult roles safely, effectively, and with an unsurpassed pleasantness of which we are all extremely proud.

Dr Peter Wilde, Head of Division

- Terms of Reference of the Review

The full review Terms of Reference and Membership are attached at Appendix A.

The Review Team set out specific areas for analysis, listed below, and this report summarises the investigations and actions taken under each topic.  During the review period, the Healthcare Commission carried out inspection to assess our compliance with the Ionising Radiation (Medical Exposure) Regulations 2000, (IR(ME)R).  A summary of the interim findings and informal feedback is also included in this document.

      - Input from all staff and assessment of clinical processes and facilities

      - Assessment of root cause analyses (RCA) undertaken on reportable incidents and on action plans agreed.

      - Analysis of all recent radiotherapy incidents to identify systemic issues, including self-assessment against Towards Safer Radiotherapy

      - Summary of the inspection of compliance with IR(ME)R regulations

      - Benchmark staffing, activity, capacity, waiting times

      - Develop detailed activity and capacity plan for 2008/09 and 2009 onwards.

- Review Reports

## Staff input into assessment of clinical processes and facilities

In order to allow staff to input into the review, an open forum or ‘Think Tank’ was organised.  Prior to the Think Tank a questionnaire was circulated asking staff to comment on their department.  The invite and questionnaire are attached at Appendices B and C.  Thirty-four responses were received, and these were themed to give four discussion topics for the evening:

- Skill mix

- Environment and equipment

- Patient pathway

- Communication

Eighty staff attended the evening event on 16 April 2008.  There were presentations on the context of the review, and then staff participated in some lively discussion around the four topics identified.  A summary of the feedback provided by staff following the event is attached at Appendix D.

In total, 115 suggestions were made about what could be done to improve the way the department runs and the staff groups interact.  These were themed and prioritised from 1 to 5, 1 being ‘quick and easy to implement’, and 5 being ‘impossible’.  Each theme has been allocated to a member of the review team to work with staff in the areas to implement.  The actions that have not already been completed will form part of a comprehensive project plan following the review.

Staff Deep in Conversation at the Think Tank, April 2008

## Assessment of root cause analyses undertaken on reportable incidents and on action plans

Between May 2007 and February 2008 four clinical incidents occurred at BHOC which were reportable to the Healthcare Commission under the IR(ME)R 2000 regulations and to the Strategic Health Authority.

In each case, a small, multidisciplinary team was brought together to undertake a root cause analysis to investigate the incident. Each team summarised their findings and then worked with the staff involved to make recommendations and develop an agreed action plan.

The table at Appendix E summarises each incident, the action taken and the progress to date.

In the detail of the investigation into U54624 and U58767, there is one common factor. In both cases, the patient asked why something in their treatment was different.  This has been dealt with by a change in the work instructions to ensure that any patient concern is investigated prior to treatment being delivered.  Human Factors Training specific to radiotherapy is also being developed.

## Analysis of recent radiotherapy incidents to identify systemic issues and assessment against

### Analysis of Clinical Incidents

Incidents in radiotherapy treatment are recorded as either Physics or Radiographic related.  An analysis of the incidents reported on Ulysses was undertaken by each Head of Department and by the Review Team, to identify any trends.

Physics incidents were analysed between July 06 and March 08, when a total of 31 were logged on Ulysses.  A total of 103 Radiographic incidents were reported between October 2007 and March 2008.  It is important to note that, in the majority of cases, these incidents were errors that were made, but picked up and rectified during the standard checking processes.

The Physics incidents were themed around calculation and transcription errors, and oversights.  From a Radiographic point of view, the incidents were clustered around the following four areas:

- Pre-Treatment, documentation of instructions/information

- Treatment data entry process and accuracy

- Document management

- Communication of intent and completion of tumour-specific information

The main areas for further action to minimise incidents are therefore in calculations and accuracy, data entry and documentation.

Calculations/Accuracy

Wherever possible, this needs to be addressed by minimising the amount of manual data entry and by adopting in-built safety checks.  The department is planning to use in-vivo dosimetry and portal imaging on all patients prior to delivery of any significant dose.  This will confirm that the expected dose will be delivered to the correct site.  Actions are underway to implement this, with a plan to complete by December 08.

In addition, the layout of the department requires improvement to enable calculations and checks to be carried out in protected quiet space away from the interruptions of the treatment floor.  An interim solution is in place in the department, and longer term this will be addressed through the BHOC’s Refresh programme.

Data entry and checking

There are checking filters in place in the department but the Review Team were interested in whether these were optimal.  The case for and against extra checking is made in the multidisciplinary report, Balancing Costs and Benefits of Checking in Radiotherapy, which recognises that checkers lower their guard if they think there will be a subsequent check.  An action arising from the review will therefore be to ensure that the checking processes followed within the department follow best-practice guidelines.

Documentation

The Clinical Governance Committee has drawn up minimum standards for documentation that should be present in the notes prior to treatment commencing.  Further work is required to ensure these documents are available, particularly where patients are transferred from external Trusts.  The Governance Committee will oversee assurance that improvements are being made and that the standards are achieved, and this will be audited through the centre’s annual audit programme.

Further detail of incident analysis is attached at Appendices F and G.

### Towards Safer Radiotherapy

In June 2006 a national multi-disciplinary working party was convened to examine the causes of errors and incidents in radiotherapy and to make recommendations for detecting errors and reducing the likelihood of their occurrence.  The working party has now published its report, Towards Safer Radiotherapy, which is endorsed by Sir Liam Donaldson, Chief Medical Officer.  The timely appearance of the report makes it possible to audit practices in BHOC against the report’s 37 recommendations.  The audit is summarised in the table at Appendix H.

Inspection of the second column of the table shows that BHOC is in compliance with the majority of the recommendations.  Several of the non-compliances are technical, and can be resolved by ensuring the Centre formalises or records some activities to demonstrate compliance.

One measure requiring further action was highlighted both by the well-publicised overdose in Glasgow and also by one of the incidents in the Centre which was reported to the Healthcare Commission.  This is the practice of normalising monitor units to a dose of one gray for all treatments, which means that radiographers on linear accelerators have to adjust these monitor units for the dose actually prescribed.  The Centre recognises the value of moving to “actual monitor units”, and is in the process of planning to implement this change.  This constitutes a significant change in practice, which must be managed very carefully to minimise the associated risks.

While the audit is largely favourable we are not complacent, and progress towards full compliance will be kept under review, both at Radiotherapy Oncology Group meetings, which reports to the Divisional Board via the BHOC Executive Group, and through specially commissioned sub-committees as required.

## Inspection of compliance with the Ionising Radiation (Medical Exposures) Regulations, 2000, (IR(ME)R) by the Healthcare Commission

In 2006, the responsibility for the inspection of organisations with respect to IR(ME)R was transferred to the Healthcare Commission and a new inspection team formed. All radiotherapy centres were advised that they could expect an inspection visit in due course.

The inspectors visited the Bristol Haematology and Oncology Centre on 17th June 2008.  Although the centre was the twenty third to be inspected, it was the first to be subjected to a revised format of inspection.

In the information letter sent to the Trust in advance of the visit the inspectors advised that they needed to be assured, not only of our compliance with the regulations but of our ability to monitor our performance, learn from incidents and continuously improve the safety of the radiotherapy service. In particular, they wished to be assured that the outcomes from the Glasgow incident had been understood in the Trust and the risks of such an incident occurring locally, minimised.

Certain documentation was requested by the inspection team ahead of the visit and this was provided.

The agenda for the inspection visit was a mixture of round table questioning and viewing of evidence together with walking through the department and questioning front line staff. Dr Jonathan Sheffield, Medical Director, represented the Trust Executive at the welcome and introductions and Ms Lindsey Scott, Director of Nursing, was present for the feedback session at the end of the day.

The feedback at the end of the day was positive but with some “housekeeping” items that required action.

A summary of the feedback is as follows:

- The inspection team were assured that the radiotherapy service complies with the regulations.

- The inspection team were assured that incidents (both errors and near misses) were escalated appropriately. They had viewed good practice in incident investigation. They considered the overview of previous incidents (see Section 7) to be impressive.

- Of the four reportable incidents, one had already been closed by them, a further two would be closed based on evidence viewed during the inspection and the fourth was on track with the action plan.

- A review of clinical audit in BHOC showed good forward planning of relevant audits with appropriate feedback at multiprofessional, monthly meetings. The inspection team were impressed that the clinical audit meetings were open to all staff at BHOC and that they were encouraged to attend.

- The inspection team identified a well established Quality Management System, accredited by the British Standards Institute (BSI). They considered that auditing of the system worked well and the service was congratulated on the evidence of rapid audit of new procedures and work instructions introduced as a result of investigation into recent incidents.

- Following discussions with front line staff, the inspection team were assured that procedures were well understood on the ground.

- The inspection team applauded the Trust for the evident open culture that encourages staff to challenge colleagues. This was endorsed by the Medical Director.

- The inspection team requires that an individual procedure is written to cover Schedules 1h and 1k of the regulations.

- The inspection team requires that procedures and work instructions are rewritten to reduce ambiguity regarding the responsibilities of duty holders.

- The inspection team requires that the age range requiring pregnancy consent be clarified and brought in line with the Trust as a whole. In addition to the written posters alerting patients to the risk of radiation if pregnant, the poster should include a pictorial message for patients.

- The inspection team requires that a method be introduced to audit if staff have read and understood emails describing changes to procedures and work instructions.

The inspection team advised that a draft report would be available to the Trust to check for accuracy within two months with the final report to follow some months later.

## Benchmark staffing, activity, capacity, waiting times

As part of the review, it was considered essential to benchmark our capacity, activity and performance with other comparable centres. In addition to developing a template of information, it was also felt necessary to visit these centres to view facilities and to see first hand how the radiotherapy service operated, with particular focus on staffing levels, activity, capacity and waiting times.

Three centres were approached for this benchmarking exercise on the advice of clinical staff and they agreed to participate. The templates completed by each of these centres were used to identify key differences which have then been evaluated and, where appropriate, included in the formulation of the overall action plan.

The visits to each centre proved extremely valuable and gave the opportunity both to learn from best practice but also to confirm the high quality of the service provided at University Hospitals Bristol. The main observations and proposed actions are set out below;

- All of the centres visited demonstrated a greater organisational focus on non-surgical oncology, both operational and strategic. At Centre 2 this was not surprising because it is a Trust specifically delivering non-surgical oncology but in the other two larger multi-speciality centres it was very noticeable that the cancer agenda was led by non-surgical oncology. This translated into active implementation of agreed strategy and high priority for service delivery issues such as equipment replacement and innovation.

- There was evidence of direct and on-going engagement with PCT’s which led to a greater understanding of the service by commissioners. This facilitated smoother LDP negotiations and more support for service developments.

- All of the centres visited had excellent facilities, a better environment than BHOC and an agreed equipment replacement methodology. In the case of Centre 1, the exterior of the hospital building was not dissimilar to that of BHOC, but internally the clinical and patient facilities had either been refurbished to a high standard or completely rebuilt.

- Despite the fact that the Radiotherapy service at UH Bristol is markedly busier than at other centres and provided in difficult accommodation, the visits team unanimously agreed that the atmosphere and staff approach within the centre was second to none.

- There were a number of interesting comparisons to be made on workforce. All of the centres visited had a higher proportion of consultants and junior medical staff than BHOC in relation to activity. In contrast, there were some differences noted in radiographer numbers and skill mix which could be explained in part by linac capacity and specialist radiotherapy provision. Also, UH Bristol seems to have far fewer A&C staff.

- As to be expected, a number of different working practices were identified during the visits which if applied at UH Bristol could lead to increased efficiency and throughput

The collated benchmarking criteria, data and comments are presented in Appendix I.

## Activity and capacity plan for 2008/09

The demand and capacity in radiotherapy is measured in the number of fractions, or exposures, which make up a course of treatment.  In 2007 the National Radiotherapy Advisory Group (NRAG) advised the government that centres would need to deliver 54,000 fractions of radiotherapy per million population (#/mp) by 2016 with each linear accelerator (linacc) working at an average rate of 8,000 fractions.  They also recommended an interim target of 40,000 #/mp by 2010/11 given capacity limitations.

The Division’s Medium Term Plan for Non-Surgical Oncology uses the NRAG recommendations to set out anticipated activity levels from 2007 to 2016.  As part of the Review, this activity data was revisited in order to provide a detailed plan for meeting the demand during the rest of 2008/09.  Although further work is underway to plan operational capacity during the opening months of Taunton centre and onwards, this report focuses on the current financial year.

In 2007/08 the radiotherapy department delivered 44,647 fractions of radiotherapy, which equates to 33,146 #/mp.  The capacity levels required to achieve the NRAG interim target are shown below.

| Year | 2007/08 | 2008/09 | 2009/10 | 2010/11 |
| --- | --- | --- | --- | --- |
| Total fractions | 44,647 | 47,699 | 32,642 | 34,629 |
| #/mp | 33,146 | 35,411 | 37,705 | 40,000 |

In 2008/09 the service has to deliver 47,699 fractions of treatment, or 35,411 #/mp in order to progress towards the 40,000 #/mp target.  A sixth linear accelerator becomes operational in June 2008.  The table below shows the capacity plan for different phases of 2008/09.

| Months | Days | Linaccs | Hours | Linacc hours/ day | #/linacc hour |  | Total # | Book at 90% | Less unplanned downtime (3%) | Total capacity (#) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Daily capacity (#s) |  |  |  |  |  |  |  |
| Apr 08-May 08 | 42 | 5 | 2x8-8 3x8-6 | 51.5 | 3.63 | 187 |  | 168 | 163 | 6,854 |
| Jun 08-Aug 08 | 64 | 6 | 6x8-6 | 58 | 3.63 | 210 |  | 189 | 184 | 11,763 |
| Sep 08-Mar 09 | 144 | 6 | 2x8-8 4x8-6 | 62 | 3.73 | 231 |  | 208 | 201 | 29,072 |
| Total | 250 |  |  |  |  |  |  |  |  | 47,689 |

The total capacity for 2008/09 expressed as #/mp is 35,404, working at 8170 fractions per linacc.

Based on the plan above, the Review team recommended to the Division that six further fixed-term Radiographer posts were required to achieve interim activity milestones, based on a total of 7.2wte staff to operate a linear accelerator, minus efficiencies.  In April 2008 the Division approved the appointment of 6wte fixed-term Radiographers to support the opening of the new linear accelerator.  The staff appointed were graduates who were not available to start until September 2008.  Longer shifts will therefore recommence from September 2008 until Taunton opens in May 2009.

- Summary of Actions Taken During the Review

The table below summarises the actions taken by the review team between April and June 2008.

|  | Work Area | Action: | Date to be completed by: | Person responsible: | Completed |
| --- | --- | --- | --- | --- | --- |
| 1 | Input from all staff and assessment of clinical processes and facilities | - Think Tank event held on 16 April 2008 - Suggestions collated into a timed action plan and leads identified - Monitor the implementation of actions - Provide feedback to staff informing them of progress | April 2008 May 2008 June 2008 May 2008 | Review Team Kate Love Review Team Jo Poulton | Completed Completed Commenced Completed and ongoing |
| 2 | Assessment of root cause analyses (RCA) undertaken on reportable incidents and on action plans agreed. | - Identify any overlapping themes and plan to redress issues. - One issue identified. Human Factors training to be planned and delivered. - Ensure individual action plans are completed and signed off. | April 2008 June 2008 July 2008 | All Jo Poulton & Helen Morgan All | Completed Commenced Commenced |
| 3a | To analyse all the recent incidents in the radiotherapy service. | - Categorise all the recent incidents and identify systemic issues - Develop action plans to deal with issues | April 2008 June 2008 | Kate Love & Alan McKenzie Kate Love & Alan McKenzie | Completed Commenced |
| 3b | To benchmark against the new standard “Towards Safer Radiotherapy” | - Identify where the service already meets the standard and where there are gaps - Develop action plans to become fully compliant | April 2008 Unsure of timescale | Alan McKenzie Alan McKenzie & Kate Love | Completed Commenced |
| 4 | To benchmark with other radiotherapy services | - Devise benchmarking criteria - Send letter inviting 3 services to benchmark - Develop benchmarking template - Make visits - Populate template and make comparisons | March 2008 March 2008 April 2008 May/June 2008 June 2008 | All Steve Falk Kate Love All Kate Love | Completed Completed Completed Completed Completed |
| 5 | To develop an activity plan to provide sufficient capacity to meet demand until Taunton opens in May 2009 | - Define the linac hours needed to meet demand until 2009. - Agree staffing levels and skill mix for those linac hours - Recruit staff (fixed term only) - Bring replacement linac into clinical use. - Review arrangements for QC checks on linacs and pilot changes. | May 2008 May 2008 May 2008 June 2008 June 2008 | Kate Love Kate Love Kate Love & Lowri Williams Alan McKenzie & Kate Love Alan McKenzie & Kate Love | Completed Completed Completed Completed Commenced |
| 8 | To respond to the minor requirements identified by the Healthcare Commission during the IR(ME)R inspection | - To develop an action plan to make the necessary adjustments to the service | June 2008 | Steve Falk, Kate Love & Alan McKenzie | Commenced |

- Final Review Action Plan

The table below shows the collated action plan arising from the review of Radiotherapy services at University Hospitals Bristol NHS Foundation Trust.

| Issue | Source/Ref. | Actions | Lead | Timescale for completion | Success Criteria |
| --- | --- | --- | --- | --- | --- |
| Ensure checking processes are robust and that accurate radiotherapy is delivered | Towards Safer Radiotherapy Incidents review | - Convert to absolute monitor units - Roll out in-vivo dosimetry - Implement changes planned as part of Refresh BHOC (calc. room etc) - Review checking processes to ensure that they are in line with best practice | Kate Love Alan McKenzie Jo Poulton Kate Love | TBC TBC Dec 2009 TBC | Reduction of incidents and near-misses related to incorrect radiotherapy being delivered. |
| Ensure that adequate checks are in place to ensure treatment of correct site | Incidents review Reportable incidents | - Review, amend and audit work instructions - Roll out Human Factors Training | Kate Love Jo Poulton Helen Morgan | Complete Sep 2008 | Reduction of incidents related to treatment of incorrect site |
| Improve availability and quality of documentation prior to decision to treat | Incidents review Think Tank | - Recruit Cancer Pathway Navigator - Audit standard criteria for documentation in the notes | Jo Poulton Paula Wilson | Jun 2008 Apr 2009 | Notes contain MDT summary, latest scans and bloods and clinic note/referral prior to treatment planning. |
| Ensure that there is an organisational focus on non-surgical oncology | Benchmarking | - Ensure UH Bristol has an integrated Cancer Strategy (including estates). This can be achieved through the work of the newly formed Cancer Board and the divisional strategy days. - Greater involvement of non-surgical oncology in Trust planning of cancer services (Cancer Board). - Develop a long term estates strategy for cancer services at UH Bristol to include future linac provision (given that two existing bunkers will be too small) | Ian Barrington Ian Barrington Ian Barrington Kate Love | Oct 2008 Aug 2008 Apr 2009 | Clear commitment to Cancer Services delivery, demonstrated through robust capital development programme and clear commissioning plan. |
| Improve engagement with commissioners | Benchmarking | - Develop regular dialogue between the commissioners and the division outside of the specific LDP process. | Jo Poulton Ian Barrington | Oct 2008 | Shared understanding of local and regional imperatives. Robust joint plans in place. |
| Improve the physical environment – space and decoration | Benchmarking Think Tank | - Achieve short-term facility improvements highlighted during radiotherapy think-tank - Deliver BHOC space plan to revised timescales supported by TEG | Ian Barrington Jo Poulton | Aug 2008 Dec 2009 | Staff and patients feel the space they use is practical, comfortable and fit for purpose. |
| Maintain good levels of staff engagement | IR(ME)R inspection, benchmarking | - Continued staff involvement in delivery and planning of services | Kate Love | Ongoing | Continuous improvement and good understanding of opportunities and pressures amongst all staff. |
| Review workforce levels and roles, particularly medical staff, administrative staff and radiographic staff | Benchmarking | - Review skill mix of radiographic staff - Continue to ensure that there is a forward plan for Consultant medical workforce and that job plans are robust and fit with the needs of the service - Review treatment planning process | Kate Love Lowri Williams Jenny Bird Kate Love | Oct 2008 Oct 2008 Dec 2008 | Patient and staff satisfaction: good use of extended roles and workforce plans in place. |
| Implement new ways of working where appropriate | Benchmarking Think Tank | - Trust Innovation Team to work with radiotherapy department during 2008/09 to ensure working practices are safe, effective and efficient | TBC | TBC | Streamlined processes. Good levels of staff and patient satisfaction. |
| Ensure procedures are clearly written and include all aspects of the IR(ME)R regulations | IR(ME)R inspection | - Rewrite individual procedures to cover Schedules 1h and 1k of the IR(ME)R regulations | Alan McKenzie | Oct 2008 | No further actions at next IR(ME)R inspection. |
| Ensure that procedures and work instructions are written so as to clearly define the responsibilities of duty holders | IR(ME)R inspection | - Rewrite key work instructions and procedures | Kate Love and Alan McKenzie | Dec 2008 | No further actions at next IR(ME)R inspection. |
| Improve regulation and information around pregnancy and consent. | IR(ME)R inspection | - Introduce new signs to raise awareness, including graphics | Kate Love | Aug 2008 | No further serious untoward incidents relating to pregnancy. |
| System required for checking staff have read and understood information informing them of changes to procedures and work instructions | IR(ME)R inspection | - Work with Trust leads and local department heads to identify how best to achieve individual and shared records of understanding of policy and competence to implement | Jenny Bird | Dec 2008 | All staff fully compliant with work instructions and protocols, with demonstrable training records. |
| Detailed future planning required. | Demand/capacity review, benchmarking | - Review Medium Term Plan assumptions to ensure demand/capacity calculations are accurate. - Prepare detailed forward plan for capacity, in conjunction with Taunton - Link forward plan to skill mix review to ensure robust workforce and financial strategy is in place | Jo Poulton Bernard Morris Kate Love Jo Poulton Kate Love Lowri Williams | Sep 2008 Oct 2008 Dec 2008 | Smooth transition through Taunton transfer; delivery of planned savings during 2009/10; achieve NRAG activity levels below reference cost. |

- Conclusion

The internal review of Radiotherapy Services at University Hospitals Bristol NHS Foundation Trust set out to establish whether the service was safe and effective.  On completion of the review, the team was assured that this was the case.

The review highlighted some areas of good practice, including:

- A culture of involvement and openness and high levels of engagement with staff

- A robust approach to reporting and learning from incidents and near-misses

- Good use of extended roles in some areas

- An extremely friendly and welcoming department

It also revealed some areas for further improvement, including:

- Need for more robust engagement with Trust Executive and local commissioners

- Need for improved facilities, environment and use of space

- Focus on ensuring robust checking mechanisms are in place to deliver the correct treatment dose

The action plan in Section 6 of this document sets out what the Review Team and Trust will do to deliver improvements and to strengthen existing areas of good practice.

The Review also highlighted the need for a local framework for ongoing horizon scanning and strategic review, which could then feed into local and Trust mechanisms for forward planning.  A new Strategy Group will be responsible for overseeing the implementation of the review action plan, reporting in to the Specialised Services Divisional Board via the BHOC Executive Group, and to the Trust Cancer Board.

Authors – The Review Team

Kate Love		Radiotherapy Services Manager

Alan McKenzie	Head of Medical Physics and MEMO

Peter Wilde		Head of Division, Specialised Services

Jenny Bird		Lead Doctor, BHOC

Steve Falk		Clinical Lead for Radiotherapy

Ian Barrington	Divisional Manager, Specialised Services

Helen Morgan	Head of Nursing, Specialised Services

Lowri Williams	HR Manager, Specialised Services

Jo Poulton		Assistant Divisional Manager, Specialised Services

June 2008

# Appendix A:  Review – rationale, terms of reference and membership of the team

DIVISION OF SPECIALISED SERVICES

INTERNAL REVIEW OF RADIOTHERAPY SERVICES WITHIN UBHT

RATIONALE

A short life review group will be established by the Division in response to the number of Ionising Radiation Medical Exposure Regulations (IRMER) reportable incidents that have occurred within the radiotherapy services since June 2007 and to concerns raised by staff over the workloads levels in the department.

TERMS OF REFERENCE

1.  To undertake a review of therapeutic radiotherapy services within UBHT, which will include:

- Input from staff working in the department and from other departments that contribute to the services e.g. medical physics, radiology, MEMO.

- Assessment of the root cause analyses undertaken for the reportable incidents and review of the agreed action plans.

- Analysis of all recent radiotherapy clinical incidents to identify any systemic issues.

- Assessment of clinical processes and facilities.

- Benchmarking of staffing (numbers, skill mix, working arrangements), activity, capacity and waiting times with other comparable centres.

- Development of a detailed activity and capacity plan for 2008/09 through review of current trends and of the assumptions made in the Medium Term Oncology Plan.

2.  To produce an Action Plan that addresses identified issues within defined

timescales, focussing on the period from now until the Taunton Centre opens in April 2009.

3.  To monitor the delivery of the Action Plan and to ensure that planned outcomes are achieved.

4.  To provide assurance to the Divisional and Trust Boards that the therapeutic radiotherapy service is safe and effective.

MEMBERSHIP OF REVIEW GROUP

| Peter Wilde, | Head of Division |
| --- | --- |
| Kate Love, | Radiotherapy Services Manager |
| Alan McKenzie, | Director of Medical Physics & Bioengineering |
| Jenny Bird, | Lead Doctor, Haematology & Oncology |
| Stephen Falk, Helen Morgan, | Lead Clinician Oncology Divisional Clinical Risk Lead |
| Lowri Williams, | Divisional HR Manager |
| Ian Barrington, | Divisional Manager |
| Jo Poulton, | Assistant Divisional Manager, Haematology & Oncology |

DATE OF FIRST MEETING

Tuesday 18th March, 2008 (Level H meeting room, BHOC)

# Appendix B:  “Think Tank” invitation

All staff involved in the radiotherapy service and Heads of Departments

BHOC

Thursday, 27 March 2008

Dear colleague

Re: Radiotherapy Think Tank,  Weds 16th April 2008, 5:30 – 9:30pm Education Centre

You will be aware that there are currently significant issues on the treatment floor including working hours, patient throughput and interaction with other departments in BHOC.  Furthermore there have been a number of highly uncharacteristic reportable incidents.  We need your help and advice to plan to improve the situation.  Novel and different approaches are particularly welcome

I hope you will be able to come to the above planned meeting where these issues can be discussed and we can see a way forward.

Food and drink will be provided

As a start I would be grateful if you could complete the attached form and return it either named or confidential as you wish to me.

Best wishes

Stephen Falk

Clinical Lead for radiotherapy

# Appendix C:  “Think Tank” questionnaire

Which 5 areas of your working practice work well?

Which five areas of your working practice hinder you in trying to do your job?

What 5 changes in working practices would make these things easier? (These changes can be small or large and need not be confined to the radiotherapy process or team if they might make a difference)

Why do you think this?

If you can think of a general solution,  how might we all make it better

# Appendix D:  Radiotherapy Review Feedack Report, April 16 2008

Highlights from Quantitative Questions

	You valued spending time discussing issues with your colleagues

	Most people found the event useful and would attend something similar again

	You would have liked a bit more time to discuss the issues that were important to you

	Some people aren’t convinced that change will happen as a result

Question		Average score

(out of 5)

- I found the event useful	3.9

- I thought the right people were here.	3.8

- I found the event informative.	3.6

- There was enough time to discuss the issues that are important to me.	3.3

- It was useful to have time to discuss issues with my colleagues.	4.0

- I feel like my views were listened to.	3.8

- I feel positive that there will be changes as a result of this evening.	3.5

- I would attend a similar event again in the future.	3.9

Highlights from Qualitative Questions

What did you find most useful about this evening’s event?

“Nice to see other people who work at BHOC and feel part of team”

“Getting everyone to agree ways to move forwards, very positive, common purpose”

“Open forum for discussion” “Chance to voice opinion… without fear of not being heard”

“Discussions with management – names to faces”

“Multi-disciplinary get-together – interesting common themes”

“The presentations” “Good use of small focus groups”

“I gained a greater understanding of other areas frustrations and concerns and that they are not just complaining”

Should anything have been done differently?

“More time to discuss issues”  “More whole group discussion”/”More time for feedback”

“V tiring at night”

More representation from other parties: registrars, ward nurses, administrative staff

Any other comments

“Better than I expected – itinerary looked rather daunting”

“Thank you”

“Many of the issues identified had been identified on previous occasions but never addressed.  I hope that this will not be the case again”

Some people commented that they would have liked time off in lieu.

# Appendix E:  Summary of Reportable Incidents and Actions Taken

| Incident | Summary | Root causes | Actions taken |
| --- | --- | --- | --- |
| Ulysses number 55803 | Lady underwent radioactive isotope treatment whilst pregnant. | Pregnancy not disclosed when checked. Potential language barrier and need for interpreter not identified with referral. | - Isotope Work instructions updated. - All endocrinologists contacted to standardise referral criteria - Referral form for isotopes changed Audit commenced - No GP’s allowed to refer patients directly - Information leaflet and pregnancy declaration form available in different languages - All actions completed, with audit completion due in June 2008 |
| Ulysses number 54624 | Patient given three fractions of radiotherapy to the incorrect side of the pelvis before it was noticed and rectified. | Differences in understanding of image inversion from patient record to simulator. | - Doctors’ induction pack developed and signed off by BHOC Exec - Progress form reviewed and updated - Protocol for bone and brain metastases written - All radiotherapy protocols on Trust DMS - New Simulator and pre-treatment check list developed and issued. - New Radiotherapy calculation room with extra work station to get calcs done away from machines provided. - Clinical Incident training for radiographers from Sue Fyffe Williams - Updated incident reporting process for serious adverse incidents implemented. - All actions above completed Human factors training being arranged for BHOC in June |
| Ulysses number 58767 | Incorrect patient treated. | Patient responded to name called in waiting area, which sounded similar to her own. Treatment given was intended for a different patient. | - New ID Work Instruction written to include ascertaining correct site. Issued and audit completed in March 2008. Results show good compliance with the new work instruction - New computer hardware provided in the treatment room to ensure electronic patient record retrieved with patient present rather than remotely - New pre-treatment work instruction regarding the use of skin marks. Issued and Audit to be done on 3rd April 2008 - Session regarding the reporting of non-conformities in the QA system planned for the radiotherapy staff meeting on 2nd April 2008. - All actions above completed |
| Ulysses number 61563 | Patient given first fraction of radiotherapy course incorrectly, More than 20% of the intended dose for that fraction was given but was corrected for over remaining fractions. | Errors in monitor unit calculations due to two systems of expressing monitor units within the department. | - Stop using IRREG to provide monitor units per fraction - Include “implementation of monitor units per fraction throughout department” on Tec Dev agenda - Compare departmental practice with “Towards Safer Radiotherapy” document - Calculation competencies added to induction and preceptorship - Audit new starters 08/09 - Update Use of Radiotherapy Assistants procedure. 4,5,6 still to be completed |

# Appendix F:  Summary of Physics Incidents

| Ulysses number | Category |
| --- | --- |
| 58200 | Calculation error, including misreading chart data |
| 57441 |  |
| 48144 |  |
| 53833 | Simple oversight |
| 48473 |  |
| 49772 |  |
| 50755 |  |
| 51099 |  |
| 58705 | Transcription error including omitting to write information on treatment plan |
| 60094 |  |
| 56010 |  |
| 56978 |  |
| 51427 |  |
| 48475 |  |
| 48502 |  |
| 56009 |  |
| 59433 | Not taking account of all available information in producing a plan or calculation |
| 55999 | Not being aware of the full range of options available on a system (such as TPS) |
| 56007 |  |
| 56008 |  |
| 56011 |  |
| 49278 | Plan delayed by oversight |
| 59196 | These are not radiotherapy physics incidents. Generally they relate to insufficient data being given to physics, so that no plan could be produced, hence error was “fail-safe”. One involved incorrect prescription according to protocol. |
| 59211 |  |
| 57273 |  |
| 57319 |  |
| 57503 |  |
| 57849 |  |
| 57850 |  |
| 55020 |  |
| 50837 |  |

# Appendix G:  Summary of Radiographic Incidents

Analysis of radiotherapy incidents from Oct 2007 – March 2008, using “Towards Safer Radiotherapy” methodology

A total of 103 radiotherapy incidents reported on the Ulysses database from 1st October 2007 until 31st March 2008 were analysed.

Each incident was coded according to the Radiotherapy Pathway Coding developed for national use in the report “Towards Safer Radiotherapy” (DH 2008).

Figure 1 shows the number of incidents for each pathway code identified.

Incidents were attributed across a range of 53 pathway codes.

32 codes had 1 incident each

10 codes had 2 incidents each

7 codes had 3 incidents each

1 code had 4 incidents

2 codes had 7 incidents each

1 code had 12 incidents

In summary the high risk categories are:

12 incidents	code10j (Pre-Treatment, documentation of instructions/ information)

7 incidents	code12f (Treatment Data Entry Process, Accuracy of data entry)

7 incidents	code19 (Document Management)

4 incidents	code 5d (Communication of Intent, completion of tumour specific information).

Action to date:

Code 10j.

A Pre-treatment checklist has been introduced to ensure that all relevant documentation is present for each patient and checked for completeness. An audit of the checklist is planned for July 2008.

Code 12f.

Some manual data entry is still required at BHOC following calculations of machine settings by the radiographers. A subgroup of the Radiation Oncology Group (ROG) is working towards minimising the calculations done by radiographers by utilising the functions on the treatment planning computer and downloading the data directly to the linac record and verify system.

In-vivo dosimetry and portal imaging will be done on all patients prior to delivery of any significant dose to confirm that the expected dose will be delivered to the correct site.

These changes require careful consideration and planning prior to implementation, including department wide training across all disciplines. Completion is expected by December 2008.

Code 19.

These incidents relate to little or wrong documentation being available when patients attend the radiotherapy department. This issue is BHOC wide and is subject to investigation by the Clinical Governance Committee.

Code 5d.

These incidents relate to poor completion of the progress form with regard to the tumour site to be treated. The numbers are small and individual incidents are fed back to the relevant staff member.

Figure 1. Numbers of radiotherapy incidents per pathway code

# Appendix H:  Summary of Audit on Compliance with Towards Safer Radiotherapy

| Recommendation | Compliance and actions |
| --- | --- |
| Radiotherapy centres should use the decision grid to classify the severity of radiotherapy errors. | To be phased in over the current year 08-09 |
| Radiotherapy centres should use: • The decision grid to classify the severity of radiotherapy errors and • The radiotherapy pathway coding system in Appendix 3.1 to identify where errors occur in a consistent manner. | To be phased in over the current year 08-09 |
| To ensure that the safe delivery of radiotherapy is maintained, each centre should formally review its skills mix and staffing levels at intervals of no more than two years and ensure these comply with national guidance. Additional reviews should be carried out during the planning of new treatment techniques or procedures and before they are introduced. | The Centre is in compliance with requirement to review regularly. However, staffing does not currently comply with national guidance. |
| Each radiotherapy centre should hold regular multidisciplinary management meetings. In addition, there should be regular multi-disciplinary meetings to discuss operational issues, including the introduction of new technologies and practices. These meetings should be informal to encourage interprofessional challenge, while respecting professional boundaries and qualifications. | The Centre is in compliance. |
| Multidisciplinary working with clear communication is essential for a safe radiotherapy department and such a culture should be actively developed. Questioning irrespective of position within the organisation should be actively encouraged. Those reporting uncertainties and errors should be given due credit for professional behaviour. | The Centre is in compliance. |
| Each department should have a fully funded, externally accredited quality management (QART) system in place. | The Centre is in compliance. |
| All procedures should be documented and subject to review every two years or whenever there are significant changes. The radiotherapy department management structure should be reviewed every two years. | The Centre is in compliance, but does not formally document the regular review of its management structure. |
| Quality policy and objectives should be reviewed at least annually and reported to a management representative appointed by the healthcare organisation. | The Centre is in compliance. |
| Training in the operation of the quality management system should be part of the mandatory induction for all staff in each radiotherapy centre. | The Centre is in compliance. |
| All centres should participate in dosimetric audit networks. | The Centre is in compliance. |
| Training records should be created and maintained for all staff involved in radiotherapy. They should be detailed and specific to particular procedures. Funding to support training should be available. | The Centre is in compliance. |
| The criteria used in the evaluation of equipment with the procurement process should include a review of both the positive and negative implications of performance specifications for patient and staff safety. | The Centre is in compliance. |
| Commissioning of radiotherapy equipment should be carried out against a written plan taking into account factors, including: • Compliance with functional specification • Clinical requirements • Statutory and regulatory requirements • Appropriate good practice guidance • Safety issues | The Centre is in compliance. |
| When new or changed treatment techniques or processes are to be introduced, a risk assessment should be undertaken and consideration given to additional verification procedures for the initial cohort of patients. | The Centre is in compliance regarding major changes but the system for highlighting smaller changes for due consideration of additional verification needs tightening. |
| All departments should have an agreed schedule of equipment quality control and planned preventative maintenance. | The Centre is in compliance. |
| The precise details of checking and verifying procedures are vital to their value. Procedures used should elicit an active response and should, as far as possible, be independent of the original method. Interruptions during radiotherapy planning and checking procedures should be minimised. | Monitor units for computer treatment plans are checked using an independent system but monitor units for unplanned patients are checked using the same charts that were used for the initial calculation. The introduction of IRREG will provide an independent check of such unplanned patients. This must await the introduction of calculating monitor units per prescribed dose rather than nominal dose (1 Gy) throughout the Centre. Such an introduction carries great risks for multiple patient harm. The introduction of IEC1217 was similarly safety critical but was successful in that no patient was harmed because of an exhaustive multidisciplinary development project to uncover all of the hidden dangers. The Centre is commissioning such a group. With the limitations of space in the Centre, interruptions during radiotherapy planning and checking are still common. |
| Correct patient identification is essential at every step. Procedures eliciting an active response from the patient must be used. The use of new technology to assist patient identification should be explored. | The Centre is in compliance with the required procedures. However, the introduction of technology such as use of photo-ID is still to be explored. |
| Each radiotherapy centre should have protocols within its quality system which define what data are to be checked by planners and prescribers along the radiotherapy pathway and how the results of these checks are to be recorded. | The Centre defines the relevant data in its quality system but needs to put in place a system for recording that these checks have been carried out for individual patients. |
| Calculations should be checked by a different entitled operator, preferably using a different method and a separate data set. Reverse checking is an example of the use of a different method. | The Centre is in compliance with regard to the allocation of checking to a different operator. Treatment plan monitor units are reverse-checked using different data sets. For independent method/data set for unplanned patients, see above. |
| Protocols should stipulate the calculation of monitor units for the actual dose to be delivered, rather than a normalised dose, to eliminate the need for additional manual calculations. It should be possible to check this as a single procedure. | The Centre is working towards this goal. Such an introduction carries great risks for multiple patient harm. The introduction of IEC1217 was similarly safety critical but was successful in that no patient was harmed because of an exhaustive multidisciplinary development project to uncover all of the hidden dangers. The Centre is commissioning such a group. |
| The use of standardised treatment protocols allows the definition of an expected range of monitor units, which provides an additional safeguard. | The Centre is compliant but the system needs to be formalised. |
| Departments should eliminate manual data transfer between computer systems. If this is not possible, then an action plan should be developed to remedy the problem and in the interim the added risk should be recognised and careful additional verification procedures established. | The Centre is largely compliant and has action plans to complete the process. |
| Each radiotherapy centre should have a clear protocol outlining the steps to be taken to ensure correct patient set-up. | The Centre is compliant. |
| Checks and verification should be performed independently by entitled operators working to clear protocols, which make explicit the individual’s responsibilities and accountability. | The Centre’s protocols need to reflect the individual responsibilities. |
| All radiotherapy centres should have protocols for on-treatment verification imaging. This should be used as a minimum at the start of a course of radiotherapy to ensure there is no gross positional error. If there is no electronic portal imaging available then film verification should be used if technically possible. | The Centre is compliant. |
| Each radiotherapy centre should have protocols for in vivo dosimetry monitoring. In vivo dosimetry should be used at the beginning of treatment for most patients. Patients should only be excluded from this procedure according to clear departmental protocols. | The Centre is in the process of introducing this. Clear protocols for excluding patients will be written once the technical feasibility for all groups of patients has been determined. The target date is December 2008. |
| Each radiotherapy centre’s protocols for in vivo dosimetry should specify action levels and the procedures to be followed for results outside the tolerance range. | The Centre is compliant. |
| Each radiotherapy centre should have an agreed policy for systematic review of patients on treatment. Concerns raised by staff must be investigated promptly. | This has not been implemented. |
| Concerns raised by patients must be taken seriously and investigated promptly. | The Centre is compliant. |
| Patient communication with staff during treatment should be facilitated. | The Centre is compliant. |
| Following a level 1 or 2 radiation incident, a systematic investigation should be conducted to identify the root causes. To prevent recurrence, the lessons learnt from root cause analysis should be disseminated locally and through a national anonymised learning system. | The Centre is compliant locally and will participate nationally when the system is set up by the Health Protection Agency. |
| Each radiotherapy centre must operate a quality system, which should ensure best practice is maintained by applying lessons learnt from radiotherapy incidents and near misses from other departments as well as in-house. | The Centre is compliant. |
| A specialty-specific voluntary system of reporting, analysis and learning from radiation incidents and near misses should be established. All radiotherapy centres should participate in this to enable national learning from safety learning. | This is being established by the Health Protection Agency and the Centre will participate. |
| Information about the error should be shared as early as possible during or after the investigation. Research into the optimal methods of feeding back lessons learnt from radiotherapy errors should be conducted. | The Centre is compliant. |
| When a clinically significant radiation incident (level 1 or 2) occurs, the patient should be informed that it has occurred and be supported in the management of any potential consequences. | The Centre is compliant. |
| When an error occurs, the staff involved should be offered appropriate support. | The Centre is compliant. |

# Appendix I:  Benchmarking data

| Criterion | BHOC | Cente 1 | Centre 2 | Centre 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Population | 1,365,715 | 1,859,948 | 1,986,685 (Some disagreement with this) | 1,330280 |  |
| Equipment Treatment Pre-treat Comp planning | 5 linacs (+ decant) 1 Synergy 1 HDR 1 Stereo 2 Sims 1 CT 1 MRI 1 U/S Mould Rm Master plan In House Radionics BrachyVision | 7 linacs 1 Orthovoltage 1 Superficial LDR Selectron (Ward based) 1 Sim 1.5 CT Advantage Sim (GE) Mould Room Eclipse x10 Brachyvision | 9 Linacs (2CBCT) 1 Cobalt HDR 2 Sims 2 CT Mould Room PLATO Pinnacle | 8 linacs (1 used as service continuity machine as per NRAG model D) 2 Varian ( 14 & 10 yrs) 5 Siemens (under 10) 1 Tomotherapy unit 1 HDR 1 SXR/DXR 1 Sim 1 CT Sessions on Neuro CT & MRI Mould Room Prosma ARPS XIO Konrad | BHOC has fewer linacs than other centres, particularly compared to Centre 3 which has a similar population. When Taunton opens there will be 7 linacs serving this population which is still less than both Centres 2 and 3 but slightly more than Centre 1 |
| Linac days per year | 246 | 253 | 253 | 250 | BHOC seems to use less linac time during bank holidays |
| Linac hours per day | 51.5 | 59.5 | 80.25 (currently up to 84.25) plus 9 hours on Cobalt | 58 – 64 (staff dependent) | BHOC provides treatment from 8:00am to 8:00pm consistently. Others are 8:00am – 6:00pm Others have multidisciplinary cover until 6:00pm. BHOC has limited cover but not after 7:00pm and no clin onc cover on site after 5:00pm |
| Total fractions in 2007/08 | 45,992 | 54,026 | 76,004 (Linacs) 7,312 (cobalt) Total 83,316 | 49765 |  |
| Fractions per linac hour (NRAG 4 – 4.5 per hour) | 3.7 | 3.7 | 3.9 (Linacs) | 3.4 | All centres provide specialist services which will reduce linac throughput. BHOC has space issues |
| Fractions per million population (NRAG 2010/11 – 40,000) | 33,676 | 29,047 | 41,937 (using pop above) We believe pop higher so probably nearer 37,870 | 38,280 | Reflects the level of linac capacity to serve population |
| Av fractions per linac per year (NRAG 2010/11 – 8,300) | 9,198 | 7,718 | 8,445 (Linacs) | 8148 (- Tomo) | BHOC is sweating the capital assests. Reflected in increased need for spare parts (Extra £100k in budget from 07/08) |
| Servicing and QA checks |  |  |  |  |  |
| Workforce WTE/Head Clinical Oncologists Consultants Senior Regs Radiotherapy Band 8 Band 7 Band 6 Band 5 Band 4 Band 2 RPU Clinical Scientists Clinical Techs Equip Engineers Trainees/non registered A&C A&C Nursing NAs | 10.31 (18) 5.00 (6) 5.40 (6) 18.31 (23) 14.47 (18) 19.00 (19) 5.61 (6) 0.00 7 (8)* 8.6 (9)* 5.5 (6)* 8 (9)* 0.2 (1) * for 6 linacs 4.00 (4) 2.00 (2) | 16.6 (17) 10.00 (10) 1.0 (1) 6.8 (7) 16.97 (23) 30.23 (32) 1.00 (1) 3.20 (4) 10.20 (11) 12.40 (13) 10.00 (10) 0.80 (1) 4.39 (5) | 12.0 (13) 6.27 (7) 50.0 (55) 20.0 (20) 2.95 (3) Band 3 9.65 (11) Band 2 2 0 (2) 6.18 (8) 0 | 21.7 (22) 8 (8) 9.3 (10) 22.47 (27) 15.2 (17) 12.0 (12) 8.0 (8) 10.7 (13) 2.1 (3) + 4.8 Rads 5.7 (6) Phys 5(5) Techs 5.6 10.34 (13) | BHOC low for clinical oncology. BHOC higher on rads to cover late shift responsibilities and part of clin onc role e.g supplementary prescribing. Total computer planning staff in Centre 3 is greater than in physics column because some managed by radiography. All centres are below Cancer Standards recommended staffing levels in physics BHOC could use A&C more effectively |
| Patient profile 2007/08 Total courses Radical Palliative/Urgent 1st Att Sub Att | 3401 1780 1621 2465 936 | 3602 | 4919 | 3680 |  |
| Total fractions | 45,992 | 54,026 | 83,316 | 49,765 |  |
| Av fractions per course | 13.5 | 15.0 | 16.9 | 13.5 | BHOC #s per course low because of limited capacity. Will increase when Taunton opens |
| Waiting time targets 2007/08 (Av no. of days) Radical Q1, Q2, Q3 Q4 Palliative Q1, Q2, Q3, Q4 Urgent Q1, Q2, Q3, Q4 | 21.7; 17.9; 18.3; 18.2 11; 8.7; 10.6; 11.3 1; 0.7; 1; 0.3 | 18; 16.3; 16.4; 16.7 8; 7.9;7.2; 7.5 0.5; 0.4; 0.1; 0.3 | 29.9, 26.8, 23.5, 25.6 10.1, 10.5, 10.0, 11.6 0.5, 0.9, 0.9, 1.3 |  | All centres have similar waiting times |
| Complex techniques 2007/08 IMRT IGRT (Synergy) Stereotactic TBI | 2 130 34 31 | 45 Yes 18 | Yes Yes No 10 (figure needs checking) | 55(14 tomo back up) 32 (Tomo since 31.10.07) 21 ( 4 single #) 8 (usually 15 – 20) | BHOC needs to identify resource to implement IMRT High level of TBI at BHOC reflects BMT programme |
| Reference Cost Index | 133 | 88 |  | 98 | Discussions at visits highlighted wide differences in costing methods. BHOC invited to be included in national group |

McKenzie, Alan et al.: Balancing Costs and Benefits of Checking in Radiotherapy, Institute of Physics and Engineering in Medicine, 2006.

The Royal College of Radiologists, Society and College of Radiographers, Institute of Physics and Engineering in Medicine, National Patient Safety Agency, British Institute of Radiology: Towards Safer Radiotherapy, London: The Royal College of Radiologists, 2008.

Radiotherapy: developing a world class service for England Report to Ministers from the National Radiotherapy Advisory Group, February 2007

2009/10 data is based on a reduced population and requires further analysis, pending decisions with Taunton about phased transfer of Somerset patients between May and August 2009.